# Documentation - I46P2_Lumipulse – Insight46 wave 2

| Categories of variables*: (may be more than one) | Simoa blood biomarkers tested on phase 1 plasma and serum samples from the LWBL/UKDRI biobank |
| --- | --- |
| Summary of work undertaken | Source data transferred to Stata file Missing data – labelled as -99. The only exception is the date variable Date_CSF_p2 where -99 in Stata gets displayed as 24/09/59, so missing data for this variable has been left as “.”. Methods section below |
| List any papers in which cleaned/derived variables have been used | - Keshavan A,Wellington H, Chen Z, et al. Concordance of CSF measures of Alzheimer’s pathology with amyloid PET status in a preclinical cohort: A comparison of Lumipulse and established immunoassays. Alzheimer’s Dement. 2021;13:e12131. https://doi.org/10.1002/dad2.12131 |

Methods

Details of the commercially available Lumipulse Aβ40, Aβ42, p-tau181 and t-tau assays are given in the kit inserts and they were all used as per the manufacturer’s instructions. All measurements were conducted on the same Lumipulse G600II platform (Fujirebio) at the Neuroimmunology Clinical Laboratory, National Hospital for Neurology and Neurosurgery, London, UK. The Lumipulse assays required 100 μL of dead volume, 50 μL/replicate for measuring Aβ42, 40 μL/replicate for measuring Aβ40, 40 μL/replicate for measuring p-tau181 and 75 μL/replicate for measuring t-tau. Measurements were undertaken in singlicate, using two reagent batches: the first 72 samples were assayed in 2019 for an interim analysis (1) and the remaining 62 samples were assayed in 2021. For the first batch, one 500 μL aliquot of neat CSF was used to measure Aβ42 and Aβ40 on a single thaw on 03/05/2019; a separate 500 μL aliquot of neat CSF was used to measure p-tau181 and t-tau on a single thaw on 06/06/2019. For the second batch, one 500 μL aliquot of neat CSF was used to measure all four analytes on 05/05/2021 and 16/06/2021 on a single thaw.

After the measurements of Aβ peptides were performed for batch 1, certified reference materials (CRM) were produced for Aβ1-42 and data have been published on the commutability of different assay platforms in measuring against these CRM (2). The values obtained by measuring Aβ42 on the Lumipulse platform were retrospectively

adjusted to incorporate standardization against the CRM, using information

kindly provided by Dr Nathalie Le Bastard (Fujirebio) (3). For batch 2, the values of the Aβ42 calibrators were already specified against the CRM.

For each of the four analytes, three quality control samples supplied with the assay kits (low, medium and high) were included on each batch run; the values of these were within the ranges specified in the kit inserts.

References

- Keshavan A,Wellington H, Chen Z, et al. Concordance of CSF measures of Alzheimer’s pathology with amyloid PET status in a preclinical cohort: A comparison of Lumipulse and established immunoassays. Alzheimer’s Dement. 2021;13:e12131. https://doi.org/10.1002/dad2.12131

2. Boulo S, Kuhlmann J, Andreasson U, et al. First amyloid beta1-42 certified reference material for re-calibrating commercial immunoassays. Alzheimer’s Dement. 2020. https://doi.org/10.1002/alz.12145.

3. Degrieck R, Dauwe M, Dekeyser F, et al. Implementation of IFCC endorsed CRMs in the LUMIPULSE® G β-AMYLOID 1-42 assay. AD/PD 2019 Poster: Fujirebio Europe NV, Ghent, Belgium; 2020.